# SiPhOG Single-Axis Angular-Rate Calibration

# Summary

SiPhOG is an IMU composed of a single-axis Fiber-Optic Gyroscope (FOG) and a single dual purpose MEMS sensor, consisting of a single triaxial rate-sensor and a single triaxial accelerometer.  To expedite the X3 calibration process, the SiPhOG will be calibrated separately from the X3, prior to assembly.  This will reduce the IC-SFvT (Scale-Factor vs Temperature) calibration from 3 separate tests to a single test (a reduction in test time from ~12 hours to ~4hours).

While a single axis calibration is insufficient to determine the sensor-axis alignment (ufogSiPhOG), a shorter, 3-axis calibration after SiPhOGs are built into an X3 makes the sensor axes observable in the X3-frame (ufogX3).  This will enable fusion of three or more FOGs into a 3D angular-rate measurement in the X3-frame, ωfogX3.

# Process Overview

This document formalizes the parameter-estimation process for the SiPhOG single-axis fiber-optic gyroscope.  The process follows the same three-step structure used for the X3:

- TC: Establishes the relationship between sensor output (in counts) to sensor temperature (in counts); this is also referred to as bias over temperature

- IC-Rate: Determines SFIC, the relationship between sensor output (in counts) to the sensor input in [deg/sec]

- IC-SFvT: Determines SFT, the effect of temperature on SFIC

However, unlike the X3 process, the linear then nonlinear calibration process is replaced with a single least-squares polynomial fit per iteration.  This is possible as the FOG is nominally aligned with the z-axis of the SiPhOG, which aligns with the rate-table rotation axis.  Manufacturing tolerances may cause the actual FOG input-axis to be misaligned with the nominal z-axis; this misalignment can be observed via a 3-axis rotation after the IMUs are installed in the X3.

Another difference from the X3 calibration process is the use of sensor counts throughout the calibration process.  Every quantity in this process stays in raw sensor counts; the sole exception is the commanded/reference rate supplied by the rate table, which is a physical truth value in degrees per second and cannot be expressed in counts.  The use of counts enables all three stages to be run in an iterative manner until the calibration parameter estimates converge.

# Sensor Model

The equation relating the 1D FOG output to its 3D input rate is:

ωfog=SFT-1⋅SFIC-1⋅ufogSiPhOGT⋅ SiPhOGR RT⋅ωinRT+ωoff+ωTempT+ωnoise

Table 0.1 lists the constituent terms in the equation.

Table 0.1: FOG Model Terms

| Term | Description | Units |
| --- | --- | --- |
| ωfog | FOG output measurement | [count] |
| SFT-1 | Sensitivity of SFIC-1 to temperature | [-] |
| SFIC-1 | FOG sensitivity to input angular-rate along sensor input axis | [count/(deg/sec)] |
| ufogSiPhOG | FOG sensor input-axis in the SiPhOG frame | [-] |
| SiPhOGR RT | Orientation of the SiPhOG-frame relative to the Rate-Table frame | [-] |
| ωinRT | Input rate in the Rate-Table frame | [deg/sec] |
| ωoff | Nominal Sensor Offset (will be included in the constant term of ωTempT) | [count] |
| ωTempT | The temperature-dependent output of the FOG | [count] |
| T | Sensor temperature | [count] |
| ωnoise | Sensor noise | [count] |

The equation above reduces to

ωfog=SFT-1⋅SFIC-1⋅ωcmdRT+ωEarthRT+ωTempT+ωnoise

by making the following assumptions:

- The misalignment vector is close to the sensor’s z-axis: ufogSiPhOG≅001T

- The sensor misalignment to the Rate-Table frame is minimal:  SiPhOGR RT≅I3

- The combination of items 1 and 2 (ufogSiPhOGT⋅ SiPhOGR RT) ensures that the input rate to the FOG is very close to the true rate

- ωoff is included in the constant term of the temperature-dependent sensor output, ωTempT; the constant term of SFIC-1 will be zero

# Calibration Stages

## Temperature Calibration (TC)

The TC process determines the temperature-dependent shift in the sensor’s data, ωTempT by placing the unit into a thermal chamber and ramping the temperature through a prescribed profile; the test is performed with the unit at rest.  The input/output relationship is:

ωout=ωin+ωTempT

Rewriting the relationship to isolate the polynomial results in:

ωTempT=ωout-ωin

As the unit is at rest in the chamber, the input rate is equivalent to Earth-rate (more specifically, the down component of the Earth-rate at the latitude of the test, φ):

ωin=ωEarth,Downφ

Complication: The sensor output is in counts while Earth-rate is in units of [deg/sec].  Prior to IC-Rate, SFIC is unknown and ωEarth cannot be accounted for.

Solution: Determine ωEarth (in counts) by following the iterative approach described below, stepping through the IC-Rate to find SFIC-1 (and SFT-1) then revisiting TC data analysis to convert ωEarth to counts by applying the sensitivity estimates to Earth rate:

ωEarth,cnts=SFT-1⋅SFIC-1⋅ωEarth

With Earth-rate removed from ωout, ωTempT can be found.

## Inertial Calibration, Rate (IC-Rate)

IC-Rate establishes the relationship between sensor counts and engineering units ([deg/sec]) by rotating the unit at known rates and relating counts to each input angular-rate setpoint.  The relationship between output and input is:

ωfog=SFT-1⋅SFIC-1⋅ufogSiPhOGT⋅ωinSiPhOG+ωoff+ωTempT+ωnoise

SFIC-1 is the sensitivity of the FOG (relating counts to rate-input in [deg/sec]) and SFT-1 is the sensitivity of SFIC-1 to temperature.  The units of SFIC-1 are countdegsec.  ufogSiPhOG is the orientation of the FOG input axis in the SiPhOG frame.  ωinSiPhOG is the SiPhOG input rate (in [deg/sec]) in the SiPhOG frame.  All other terms are in counts.

Rewriting this in a matrix format (A⋅x=b) enables the sensor parameters to be estimated by inverting the design matrix, A, and solving for x.

ufogSiPhOGT⋅ωinSiPhOG1⋮⋮⋅SFT-1⋅SFIC-1ωoff=ωfog-ωTempT⋮

However, as mentioned above, ωoff is included in the constant term of ωTempT and the constant term of SFIC-1 is set to zero.  This resolves the identifiability problem of which polynomial “owns” the constant terms.  The formulation can be rewritten as:

SFT-1⋅ωrtωrt2⋯ωrt5⋮⋮⋮⋮⋅a1a2⋮a5=ωfog-ωTempT⋮

where ωrt is the rate table rotation rate (rationale described in next section):

ωrt=ufogSiPhOGT⋅ωinSiPhOG

Note: in the formulation above, the rate command design matrix is pre-multiplied by SFT-1 (a scalar) to compensate for the temperature dependency of SFIC-1.  Alternately, the entries in the measurement vector, b,  can be divided by SFIC-1.

When calibrating the SiPhOG directly, it is only spun about its z-axis so ufogSiPhOG cannot be determined during this test.  However, ufogSiPhOG should be close to zSiPhOG so ufogSiPhOGT⋅ωinSiPhOG~ωRTSiPhOG, where ωRTSiPhOG is the rate-table motion.  In this configuration, the system is least sensitive to any misalignment of ufogSiPhOG and zSiPhOG.

Note: In previous implementations with the X3, this is broken into a linear and a nonlinear solution for SFIC-1 (SFT-1 is treated as 1.0 until determined in the next step).  First, values for SFIC-1 and ωoff are estimated over a small range of input rates (-100 dps≤ωinSiPhOG≤100[dps]) then SFIC is found by fitting a polynomial between the x-axis data (x=ωfog-ωoff-ωTempT) and the y-axis data (y=ufogSiPhOGT⋅ωinSiPhOG=ωRTSiPhOG); SFT set to 1.0 for the first pass through the estimation.

SFT⋅SFIC⋅ωfog-ωoff-ωTempT=ufogSiPhOGT⋅ωinSiPhOG

## Inertial Calibration, Scale Factor vs Temperature (IC-SFvT)

IC-SFvT establishes the dependency of SFIC-1 on temperature.  The SiPhOG is spun at two rates (+200 [dps] and -200 [dps]) while the temperature of the system is changed.  The sensor output is recorded along with the SiPhOG temperature.  Note: in the following formulation for SFT-1,  the noise term is transformed into a noise ratio.

SFT-1=ωfog-ωTempTSFIC-1⋅ufogSiPhOGT⋅ωinSiPhOG+ε

Where the rate terms in the denominator can be replaced by the rate-table rate, ωrt:

ωrt=ufogSiPhOGT⋅ωinSiPhOG

The process described above is repeated until the parameters converge.

Once SiPhOGs are assembled into the X3 (three-axis IMU), the unit is subjected to rotations about all three SiPhOG axes and ufogX3 is determined.

## Data Processing

FOG Calibration is performed in three stages (TC, IC-Rate, and IC-SFvT), which are described later in this section.  Data is processed in an iterative manner following data collection, as outlined below.

Step 1: TC data is processed (without Earth-rate correction) to determine 
ωTempT.  IC-Rate data is used to find SFIC-1. IC-SFvT is not run.

Step 2: TC data is re-processed (with Earth-rate correction using the newly found SFIC-1).  IC-Rate data is re-analyzed (using the updated ωTempT) and IC-SFvT data is processed to find SFT-1.

Step 3, 4, 5, …: Data is re-analyzed with all three dataset processed on every pass, each using the most recently available estimate of the calibration parameters, until convergence.

Convergence is checked on the fitted response curves, not the raw polynomial coefficients — a poorly constrained high-order coefficient can swing by a large relative amount between iterations while the curve it produces barely moves, which would make a coefficient-based tolerance depend on which term happens to be noisy rather than on how much the calibration itself changed. For each block, its curve is evaluated on a grid spanning that block's observed input range, and the relative change is:

ΔX = max | curvenew − curveold | / ( max | curveold | + ε )

Δ = max( ΔTC, ΔIC-Rate, ΔIC-SFvT )

The process stops when Δ falls below a chosen tolerance, or when a maximum iteration count is reached — whichever comes first — and reports which of the two conditions triggered the stop, so a run that hit the iteration cap without truly converging is visible rather than silently accepted.

# Calibrated Output

The output of the SiPhOG, following calibration, will be the best estimate of the input rate, projected onto the FOG’s input axis:

ωest≅ωproj=ufogSiPhOGT⋅ωinSiPhOG

=SFT⋅SFIC⋅ωfog-ωTempT-SFT⋅SFIC⋅ωnoise

Ignoring the noise term, the output is:
ωest=SFT⋅SFIC⋅ωfog-ωTempT

# Estimate of ufogX3 and Fused Rate Output

This section is a placeholder for estimating ufogX3 when the SiPhOG is built into an X3…

Nominally.  Although the number of MEMS has no bearing on the calibration process described here.

Note, once assembled, the input axes estimates are relative to the X3-frame.